Clover 10 handheld video magnifier

HC124

Please retain these instructions for future reference. These instructions are also available in other formats.

Contents

Special warning	2

General description	2

Items supplied with the Clover 10	2

Orientation	2

Front	3

Back	3

Left edge	3

Right edge	3

Top edge	3

Using the product	3

Battery charging	3

Power LED	4

Stand	4

Operation	4

Power Button	4

Beep Type	4

Beep Control	4

Camera Switch	5

Magnification	5

Colour Mode	5

Brightness	6

Lighting	6

Image Freeze	6

Image Storage	6

Image Playback	7

Line & Mask	7

Find	7

Real-time Navigation	7

Memory	8

HDMI Output	8

Auto Power Saving	8

Restore	8

Hardware Reset	8

Technical specification	8

Hints and Tips	8

Troubleshooting	8

The unit does not turn on	8

Black screen or White screen	9

The image does not change	9

Smudges or blemishes on the screen	9

How to contact RNIB	9

Terms and conditions of sale	9

# Special warning

This product is returnable within 14 days of receipt in an “as new” condition in original, undamaged packaging.

Before using Clover 10 for the first time, you must charge its battery. Refer to Battery Charging section for more details.

Always operate the unit in a place where the temperature is between 10°and 35°C (50°and 95°F).

Keep the unit away from wet areas, rain, and other sources of liquids. Do not submerge the unit.

Do not attempt to repair the unit. Any unauthorised attempt will void the product warranty. Please contact your distributor.

Use the supplied power adapter and battery.

# General description

Clover 10 is a portable 10 inch video magnifier with excellent full HD image quality, and it assures a real-time crisp image while moving or changing magnification continuously from 2.5X to 18X. The integrated dual camera system provides three viewing ways to fulfil the daily use.

## Items supplied with the Clover 10

Power adaptor

HDMI cable

Cleaning cloth

Carrying case

## Orientation

Place the Cover 10 in front of you so that the screen is facing towards you and the buttons are running along the bottom edge.

### Front

Along the top edge of the front and in the centre, is a small circular camera lens. Just below this is the LCD screen.

Below the LCD screen and running along the bottom edge are five buttons. From left to right these buttons are:

Enhanced colour

Zoom out

Joystick

Zoom in

True colour

### Back

Turn the Clover 10 over so that the LCD screen is facing down.

Just below the top edge in the centre is the rear camera, and to the left and right of this are two recessed LED lights.

Running along the bottom and left and right edges is the fold-out stand.

### Left edge

Along the left edge are four buttons and sockets. From top to bottom these are:

Power button

Power port

HDMI port

Locate button (located at the bottom)

### Right edge

On the right hand edge, at the bottom, is a single button which is the freeze button.

### Top edge

Along the top edge, towards the left hand side, is the camera slide button. Just to the right of this is the USB port.

# Using the product

## Battery charging

The power adaptor comes assembled. Connect the power cable to the unit’s power port.

Plug the power adapter into an AC outlet and start charging.

### Power LED

Charging when the unit is off, the power LED will be green, and then auto turn off when fully charged. A full Charge will take about five hours, and will support the unit working for about 3 hours.

Charging when the unit is on, the power LED will be green.

When the unit is on without charging, the power LED will be green when the battery level is high and will be red when the battery is dying. As long as the red light starts to flash, it is a signal to charge the battery.

You can also check the battery status by pressing the power button.

## Stand

With integrated fold-away writing stand at an ergonomic viewing angle, Clover 10 can satisfy your desire to read more and long time in a comfortable position. There’s also enough space under the camera to work the pen. Unfold the stand by gently pulling the two legs on the left and right hand side out.

The rear camera is designed to auto rotate according to the status of the stand, so that both the near view and far view are supported on this unit.

## Operation

### Power Button

Press and hold the power button for 5 seconds to turn on/off the unit.

When the unit is on, press the power button to check the current status of battery capacity, freeze mode, focus status, and beep.

Press other buttons when holding the power button will make combo functions. Refer to the following instructions for more details.

### Beep Type

Unit on/off beep

Button beep

Threshold beep

Low battery beep

### Beep Control

Press and hold the power button and press the true colour button to turn on/off the beep.

### Camera Switch

Slide the camera switch left to open the rear camera for near view and far view.

Slide the camera switch right to open the front camera for mirror view.

### Magnification

Press the zoom in button to increase magnification, and hold the zoom in button for continuous increase.

Press the zoom out button to decrease magnification, and hold the zoom out button for continuous decrease.

Note: Magnification range when standing on table: back/front camera: 2.5x to 18x.

### Colour Mode

A colour mode is a predefined colour combination. Full-colour mode is suitable for viewing pictures or maps.

Enhanced-contrast mode is combined with different foreground and background colours which can help to read more clearly.

Press the enhanced colour button to go through the enhanced-colour combination list.

Press the true colour button to return to the full-colour mode.

Hold the enhanced colour button for 5 seconds to switch between the simplified enhanced colour combination list and the integrated enhanced colour combination list (factory default).

The simplified enhanced colour combination list (5 modes):

White on black

Black on white

Yellow on black

White on blue

Yellow on blue

The integrated enhanced colour combination list (18 modes)

White on black

Black on white

Yellow on black

Black on yellow

White on blue

Blue on white

Yellow on blue

Blue on yellow

Green on black

Black on green

Violet on black

Black on violet

Blue on black

Black on blue

White on green

Green on white

White on red

Red on black

### Brightness

Brightness of the screen can be adjusted in both full-colour mode and enhanced-contrast mode.

Press and hold the power button and press the zoom in button to increase the brightness. Factory default value will be highlighted green.

Press and hold the power button and press the zoom out button to decrease the brightness.

### Lighting

Press and hold the power button and press the freeze button to turn on/off the LED lighting manually.

### Image Freeze

Freezing an image on the LCD screen can be useful for further reading. At this status, magnification as well as colour change is still allowed for a better experience.

Press freeze button to lock the active image.

Press freeze button again to unlock.

For the locked image with increased magnification, you can use the joystick to navigate around freely.

### Image Storage

Long press the freeze button for 3 seconds to store the current image. You can save up to 18 images.

Note: When the saved images reach to the maximum value (18 images, memory full icon appears), you will need to delete some pictures first in playback mode before the next successful saving.

### Image Playback

To enter the playback mode, long press the true colour button for 3 seconds.

Slide the joystick to select the target image from the nine grids thumbnail image gallery. Press the true colour button to view the target image. You can view and navigate the image with the same operations described in the image freeze section.

Press the freeze button to return to the nine grids thumbnail.

When in nine grids thumbnail, press the locate button to select the target image which needs to be deleted. Press the locate button again to confirm, or any other buttons to cancel.

When in nine grids thumbnail, hold the locate button for 5 seconds to delete all the images. Press the locate button again to confirm, or any other buttons to cancel.

Long press the true colour button for 3 seconds to exit the playback mode.

### Line & Mask

Press the locate button to switch among different reading line & mask modes.

One horizontal line

One vertical line

Horizontal mask

Normal, no line & mask

Slide the joystick to change the position of reading line or the distance between the masks. The real-time navigation is unusable in line & mask mode.

Press the power button to quickly exit the line & mask modes.

### Find

Long press the locate button for 2 seconds and hold not to release it, a location box will appear.

Move and change the content in the location box, aiming it at the part you want to view, and then release the locate button, the image will return back to the previous magnification with the refreshed content.

For example, you are viewing one part of a map with a certain magnification. Then you want to see another part without changing your favourable magnification. Then the function would be convenient and useful for you.

### Real-time Navigation

In the state of amplification, users can slide the joystick when in the real-time video mode to navigate the image for seeing more without moving the unit.

Press the zoom in and zoom out buttons at the same time to help you re-centre the image (red dot appears).

### Memory

The unit will memorise automatically the most recent settings such as colour mode, magnification, beep when power is off normally and apply them when the unit is on next time.

### HDMI Output

Connect the unit to your external display device with HDMI cable for a better and larger viewing.

### Auto Power Saving

If no operation and no move for over 3 minutes, the unit will turn off automatically.

Press and hold the power button and click the enhanced colour button  to turn on/off the auto power saving function manually (factory default is on).

### Restore

Long press the freeze button and the locate button at the same time for 5 seconds to reset to the default factory settings. Auto power off when reset is done.

### Hardware Reset

When the unit is hung or can’t be turned off normally, press the power and freeze button and true colour button at the same time to trigger the hardware reset.

# Technical specification

Display: 10 inch wide LCD, HD screen

Camera: Auto Focus, HD

Size: <= 245 x 200 x 22mm

Weight: <=680 grams

Colour Mode: True colour, 5 Simplified enhanced colour modes; 18 integrated enhanced colour modes

Magnification: 2.5x – 18x

Battery: Integrated rechargeable Li battery; 3-hour continuous use after 5-hour charge

Power Adapter: Input: 110-240V; Output: 5V/3A

# Hints and Tips

## Troubleshooting

### The unit does not turn on

Connect to the power adapter

Charge the battery

Hardware Reset

### Black screen or White screen

If the unit is sitting on a flat pure surface, the screen may be black or white. Lift the unit off to ensure change happens

Adjust the magnification

Adjust the colour mode

Restart the unit

Hardware Reset

### The image does not change

If the unit enter the freeze mode or playback mode, the image will not change

Press power button to check the current status

Restart the unit

Hardware Reset

### Smudges or blemishes on the screen

Adjust brightness

Clean the LCD screen and camera

# How to contact RNIB

Phone number: 0303 123 9999

Textphone: 0845 758 5691

Web address: rnib.org.uk/shop

Postal address: RNIB Peterborough, Midgate House, Peterborough PE1 1TN

Email: helpline@rnib.org.uk

RNIB Technology Team

Telephone 0207 391 2280

Email tfl@rnib.org.uk

Email for international customers exports@rnib.org.uk

## Terms and conditions of sale

This product is guaranteed from manufacturing faults for 24 months from the date of purchase.  If you have any issues with the product and you did not purchase directly from RNIB then please contact your retailer in the first instance.

For all returns and repairs contact RNIB first to get a returns authorisation number to help us deal efficiently with your product return.

You can request full terms and conditions from RNIB or view them online.

Registered Charity No. 226227

This product is CE marked and fully complies with all applicable EU legislation.

At the end of the product’s life you can return it to the supplier or manufacturer for free of charge disposal.  This is to comply with ‘The Waste Electrical and Electronic Equipment Directive’ (WEEE).

Date: June 2017

© RNIB